Tenma DC Electronic Load – Guide and application notes

This document is written to give context and guidance in using the Tenma (or more broadly any) electronic load for use in PCB test and automation.

Warning!

Electronic loads are not as simple to use, and cannot fully replace dummy resistor loads etc, in all situations. Electronic loads have control loops fine tuning the parameters to keep the desired output (CV,CC,CR,CW). These control loops can fight against the switchmode/LED drivers, causing oscillations and potentially resulting in damage to the PCB / electronic load / or both!

-----!

Operating modes

CV	-	Keeps the Voltage constant.

CC	-	Keeps the Current constant. Normally the best mode to replace a load resistor when trying to apply a consistent current draw to an output.

CR	-	Keeps the Resistance constant.

CW	-	Keeps the Power constant.

Benefits of using an electronic load

There are a couple key advantages to using an electronic load over power resistors.

- Precise tuning of loads – easily tune to mA

- Multiple values can be switched in to profile the channel in more detail without user interaction

- Not voltage dependant – Resistor loads have to be calculated based on a fixed voltage

Space – several load values and associated heatsinks soon take up considerable shelving space and take time to find/setup on the test bench

- Voltage, current, and power measurement built in

Weaknesses

- Cost – larger upfront cost ~£200.

- Only 1 channel can be tested simultaneously

- Risk of oscillations and damage to DUT and Load if not considered and configured correctly

Connection and communications

The load can be controlled from the front panel or any of the rear comms ports (USB, RS-232, Ethernet). Message structure is the same between each. The easiest connection is USB which presents a COM port to the PC without additional adaptors. This can then be integrated into the F&F Test Script tool.

The baud rate can be set on the unit and this needs to match in the terminal or test tool. Advise setting to the maximum of 115200kbps.

Terminal settings:

Baud rate: set to 115200

Termination: LF/CR ?????

To test the comms are successful, send the command IDN? This should return the product information. Full set of commands can be found in the Appendix.

OCP/OVP

The use of over-current and over-voltage protection is highly advised.

Example

The below example tests the average load/voltage, peak load, and overload capability of a USB3 port on the Discover Base control PCB. Comparing this to the current test procedure whereby only the average load is tested, the electronic load allows the test procedure to quickly check average, peak, and OL conditions, while monitoring the voltage sag from the 5V USB supply and confirming the port is being loaded as required.

Operator interaction is only required to swap to the next port. Monitoring the voltage in the test tool, the port test could start automatically when the 5v appears.

Full set of commands can be found in the Appendix.

276-020-AS-03	Discover Base USB and Control

USB3 ports

| Command | Response | Description | Notes |
| --- | --- | --- | --- |
| :INPut OFF |  | Turn OFF | Start from known good state |
| :FUNC CURR |  | Configure for CC |  |
| :CURRent 0.5A |  | Configure for CC 500mA | Average load |
| :INPut ON |  | Turn ON |  |
| :MEASure:VOLTage? | >4.845V | Measure Voltage | Check voltage sag >4.8V |
| :MEASure:CURRent? | >0.502A | Measure Current | Check current is ~500mA |
| :CURRent 0.9A |  | Configure for CC 900mA | High load |
| :MEASure:VOLTage? | >4.721V | Measure Voltage | Check voltage sag >4.6V |
| MEASure:CURRent? | >0.897A | Measure Current | Check current is ~900mA |
| :CURRent 1.2A |  | Configure for CC 1200mA | Trip load |
| :MEASure:VOLTage? | >0.0000V | Measure Voltage | Check voltage is 0V |
| MEASure:CURRent? | >0.000A | Measure Current | Check current is 0mA |
| :INPut OFF |  | Turn OFF | Resets USB port |
| :CURRent 0.5A |  | Set CC 500mA | Average load again |
| :INPut ON |  | Turn ON |  |
| :MEASure:VOLTage? | >4.845V | Measure Voltage | Check voltage sag >4.8V |
| MEASure:CURRent? | >0.501A | Measure Current | Check current is ~500mA |
| :INPut OFF |  | Turn OFF | Finished |

| Communication Commands with Computer V2.10 (table sourced from KEL103 manual and resources) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| S/No. | Command Modes |  | Command Interpretation | Command Examples |  |
|  | Commands | Parameter Description |  | Setup | Query |
| 1 | IDN？ | no | Return product information | no | *IDN? Return product information |
| 2 | *SAV | <NR1> 1-100 | Store to unit | The current value is saved in 20 | no |
| 3 | *RCL | <NR1> 1-100 | Recall the storage unit | *RCL 20 Recall 20 units | no |
| 4 | *TRG | no | Simulate an external trigger command, only valid in the pulse mode and the flip mode. | *TRG Simulate trigger once | no |
| 5 | :SYSTem :SYST | BEEP¦BAUD | Set system parameters such as buzzer, baud rate setting query etc. | :SYSTem:BEEP ON Beep ON | :SYSTem:BEEP? >ON Query the beep ON |
| 6 | :STATus? :STAT? | Baud rate: 0,9600 1,19200 2,38400 3,57600 4,115200 | Query the device status The first byte is the buzzer status and the second byte is the baud rate; other bytes are to be determined. | no | :STATus? >0,4,0,0,0,0 |
| 7 | :INPut | <Boolean>¦OFF¦ON | Set the query of the device ON/OFF | :INPut ON Switch on the device |  |
| 8 | :VOLTage :VOLT | <NR2>MAX¦MIN | Set CV voltage and query CV voltage | :VOLTage 20V Set the CV voltage as 20V | :VOLTage? >20V The CV voltage is 20V |
| 9 | :VOLTage:UPPer？ :VOLT:LOWer？ | no | Query the max/min voltage of the device | no | :VOLTage:UPPer？ >150V Return the max voltage of the device |
| 10 | :CURRent :CURR | <NR2>MAX¦MIN | Set the CC current and query it | :CURRent 2A Set the CC voltage as 2A | :CURRent? >2A The CC current is 2A |
| 11 | :CURRent:UPPer？ :CURR:LOWer？ | no | Query the max/min current | no | :CURRent:UPPer？ >30A Return the max current of the device |
| 12 | :RESistance :RES | <NR2>MAX¦MIN | Set the CR resistor and query it | :RESistance 20OHM Set the CR resistor as 20Ω | :RESistance? >20OOHM The CR resistor is 20 Ω |
| 13 | :RESistance:UPPer？ :RES:LOWer？ | no | Query the max/min resistor of the device | no | :RESistance:UPPer >7500OOHM Return the max resistor of the device |
| 14 | :POWer :POW | <NR2>MAX¦MIN | Set the CW power and query it | :POWer 20W Set the CW power as 20W | :POWer? >20W The CW power is 20W |
| 15 | :POWer:UPPer？ :POW:LOWer？ | no | Query the max/min power of the device | no | :POWer:UPPer？ >300W Return the max power of the device |
| 16 | :FUNCtion :FUNC | VOLT¦CURR¦RES¦POW¦SHORT Define the functions: voltage, current, resistor and power | Only can switch CV, CC, CR, CW Can query CV, CC, CR, CW, that in continuous mode, pulse, flip, battery and all the other modes. | :FUNCtion VOLT Set the constant voltage mode | :FUNCtion？ >VOLT The current mode is constant voltage |
| 17 | :MEASure:CURRent | <NR2> | Return the load current. Unit A | no | :MEASure:CURRent? >0.789A |
| 18 | :MEASure:VOLTage | <NR2> | Return the load voltage. Unit V | no | :MEASure:VOLTage? >1.4999V |
| 19 | :MEASure:POWer | <NR2> | Return the load power. Unit W | no | :MEASure:POWer? >1.1968W |
| 20 | :LIST |  | Output all steps in order | :LIST 2， 3A,2,1A,0.1A/uS,5S,2A,0.1A/uS ,5S,3 This command means CC range 3A, LIST step is 2 steps; the current of the first step is 1A and the slope is 0.1A/uS; the time is 5s; the current of the second step is 2A, the slope is 0.1A/uS, the time is 5s. Repeatedly run 3 times and saved in LIST 2. | no |
| 21 | :RCL:LIST | 1-7 | Recall/query the LIST unit Recall the query unit before query, or an unknown condition occurs. | :RCL:LIST 2 Recall LIST 2 | :RCL:LIST？ >3A,2,1A,0.1A/uS,5S,2A ,0.1A/uS,5S,3 Query LIST 2 |
| 22 | :OCP |  | Output all steps in OCP mode | :OCP 10V,5S,3A, 0.1A,0.1A,0.2S,2A,2V,1.5A,0.9 A,5 This command means: VON voltage is 10V, VON delays 5s, current range is 3A. Initial current is 0.1A, stepping current 0.1A, stepping delay 0.2s, cutoff current 2A, OCP voltage 2V, max overcurrent 1.5A and min 0.9A, which save in LIST 5. | no |
| 23 | :RCL:OCP | 1-10 | Recall/query the OCP's unit. Recall the query unit before query and then use the query. | :RCL:OCP 2 Recall the OCP mode and save in LIST 2. | :RCL:OCP？ >10V,5S,3A, 0.1A,0.1A,0.2S,2A,2V,1 .5A,0.9A,2 Query the value of OCP mode in LIST 2 |
| 24 | :OPP |  | Output all steps in OPP mode | :OPP 10V,5S,5A, 0.1W,1W,1S,12W,7V,6.5W,5.6W,5 This command means: VON voltage is 10V, VON delay 5s, current range 5A, intial power 0.1W, stepping power 1W, stepping delay 1s, cutoff power 12W, OPP voltage 7V, the max overpower 6.5W and min 5.6W, which save in LIST 5. | no |
| 25 | :RCL:OPP | 1-10 | Recall/query OPP's unit. Recall the query unit before query and then use the query. | :RCL:OPP 3 Recall the OPP mode and save in LIST 3. | :RCL:OPP？ >10V,5S,5A, 0.1W,1W,1S,12W,7V,6.5W ,5.6W,3 Query LIST 3 |
| 26 | :BATTery :BATT |  | Output all steps according to the battery mode | :BATT 2,30A,7A, 35V,11AH,50M This command means: setting the battery range 30A, discharge current 7A, discharge cutoff voltage 35V, discharge cutoff capacity 11AH, and discharge time 50M, which save in LIST 2. | no |
| 27 | :RCL:BATTery :RCL:BATT | 1-10 | Recall/query the battery unit. Recall the query unit before query and then use the query. | :RCL:BATT 5 Recall the battery storage LIST 5 | :RCL:BATT？ >30A,7A,35V,11AH,5M Query LIST 5 |
| 28 | :BATTery:TIM |  | Query the time in the battery mode | no | :BATTery:TIM >5.00M |
| 29 | :BATTery:CAP |  | Query the capacity in battery mode | no | :BATTery:CAP >5.00AH |
| 30 | :DYNamic | <NR1> 1-6 | Query and set the dynamic test mode: 1 refers to dynamic CV, 2 refers to dynamic CC, 3 dynamic CR, 4 dynamic CW, 5 indicates pulse mode and 6 flip mode. | :DYN 1,1.0001V,2.0003V,1.101HZ,50. 0001% dynamic CV | :DYN? >1,1.0001V,2.0003V,1.1 01HZ,50.0001% Return the value of the dynamic mode; set before using this command. |

Note:

- The unit voltage of all commands is V, current A, the current slope A/uS, the resistance OHM, power W,time S or M ( S-second, M-minute ), capacity AH and percentage %.

- All commands can be abbreviated in SCPI format, for example, DYNamic? and :DYN? are the same command.